Stepwise Process for Improving the Quality of Rapid Diagnostic Testing (SPI-RDT)

Users’ Guide for Site Assessment Using the SPI-RDT Checklist

2026

Contents

Background	2

Purpose of the Users’ Guide for the SPI-RDT Checklist	3

Steps and Requirements of the Assessment	3

Pre-Assessment Steps	3

Assessment	4

Post-Assessment	4

Tips on How to Conduct the Assessment	4

Structure of the SPI-RDT Checklist	6

Part A.  Characteristics of the facility and testing site assessed	6

Part B. SPI-RDT Checklist	7

Part C. Scoring	22

Part D. Site assessment Report and Recommendations	23

# Background

Global guidance is increasingly promoting integrated testing services and coordinated quality management approaches at the point-of-care (POC), particularly as countries expand access to rapid diagnostic testing across multiple diseases. Within Ministry of Health (MoH) tiered laboratory networks, diagnosis and treatment services are progressively shifting to lower levels of the health system, supported by simple, rapid tests performed at or near the patient.  Rapid diagnostic tests are widely used at many point-of-care testing sites to expand access to high-quality healthcare services and improve patient outcomes. As rapid testing grows across multiple disease programs, disease-specific tools and parallel systems can create redundancies and fragmentation. In addition, many programs do not have quality assurance systems fully in place. Global guidance increasingly emphasizes integrated testing services and coordinated quality management approaches at the POC, especially as countries expand rapid diagnostic testing for multiple diseases. An integrated approach helps maximize existing resources, reduces fragmentation, and improves long-term sustainability. However, this expansion must be accompanied by robust quality assurance systems to ensure test results remain accurate, timely, and reliable for patient management.

Strong POC testing sites require essential supporting systems and tools: up-to-date standard operating procedures for all testing processes; appropriate safety equipment and adherence to safety practices; reliable procurement and inventory management; routine quality control and external quality assessment (EQA) for all tests conducted; a trained and competent workforce; adequate on-site supervision; and effective referral pathways for confirmatory or specialized testing. Strengthening these cross-cutting elements supports quality across all rapid testing services delivered at the same site.

The Updated Checklist: Stepwise Process for Improving the Quality of Rapid Diagnostic Testing (SPI-RDT)

To support users in assessing and strengthening quality at POC testing sites, an updated Stepwise Process for Improving the Quality of Rapid Diagnostic Testing (SPI-RDT) checklist has been developed. This new checklist replaces the previous Stepwise Process for Improving the Quality of HIV Rapid and Recency Testing (SPI-RRT), which was originally designed to establish minimum quality standards and quality assurance (QA) practices for decentralized HIV rapid and recency testing sites and to support continuous quality improvement (CQI).

As the testing landscape evolves, there is an increasing need for a checklist that can be applied across multiple disease programs. Building on the success of the original checklist, the revised SPI-RDT checklist was updated to be pathogen-agnostic while retaining the proven strengths of the SPI-RRT approach, including its stepwise methodology and focus on CQI, while expanding it to enable pathogen-agnostic assessment of any rapid diagnostic testing services delivered at the same testing site. Key updates include:

- Streamlining the checklist from eight sections to five to simplify use and improve efficiency.

- Updating the testing section to be pathogen-agnostic, so one checklist can be used to assess multiple testing programs operating within the same site, using the same set of elements.

- Clarifying that the checklist assesses the overall testing program (systems, processes, and quality practices), rather than evaluating an individual test.

- Reorganizing content, so pre-test, test, and post-test elements are integrated throughout the sections as cross-cutting components of quality.

- Adding an optional specimen referral section and revising point values to align scoring with the updated structure and quality expectations.

With these updates, the SPI-RDT checklist is designed to provide a unified and efficient approach to quality improvement for high-quality test results in decentralized settings across multiple disease programs. We anticipate that one site visit and a single checklist could cover all pathogens, saving time and reducing duplication. All testing programs would follow the same quality improvement framework, and assessors could be trained once to support any rapid diagnostic testing program. This consistency is expected to make it easier to identify shared gaps that affect multiple tests and allow the approach to adapt quickly as new diagnostics are introduced. Considering all testing programs together should support better planning and resource use, a strong quality framework would be in place to respond to emerging threats or new technologies.

Scope and Application

The updated checklist is intended for use at the testing-site level where POC testing is conducted. Where multiple testing sites operate within the same facility, each site is treated as an independent assessment unit and should be assessed separately.

It is important to note that the checklist is a flexible template that can be adapted by countries to meet their specific needs. Countries are encouraged to utilize the SPI-RDT tool in various facility types, including laboratories, where RDTs may also be performed. This adaptability allows the checklist to effectively address quality gaps in RDT use across different testing environments, ensuring that it remains relevant and supportive of quality improvement initiatives in both POC and laboratory settings.

By customizing the checklist, countries can enhance its applicability and maximize its potential to improve the quality of rapid diagnostic testing services delivered in their health systems

## Purpose of the Users’ Guide for the SPI-RDT Checklist

The Users’ guide provides instructions for implementing the checklist in an accurate and standardized manner. It explains the intent of each question and how to assess it, including what data to collect and how to verify evidence through document review, interviews, and/or observation. Examples provided in Section IV, “Tips on How to Conduct the Assessment” are used to illustrate the methods described in that section.

Some checklist items may be outside the purview of the testing site personnel; in these cases, site supervisors and/or higher levels with MoH may need to address broader management or systems level issues. Aggregated assessment results can also support advocacy, program planning, and policy decision-making at MoH level.

The checklist, particularly the testing section, is designed to be flexible across different algorithms and pathogens. Assessors should be competent in all checklist areas, apply evidence-based scoring, and calculate total possible and achieved scores correctly.

## Steps and Requirements of the Assessment

Notes:

- When conducting assessments, they should be carried out as an integrated process led by the MoH team collaboratively (not in parallel or independently).

- Two assessors per site visit are preferred to improve objectivity when conducting an assessment. However, if resources are limited, a single assessor may conduct the assessment.

- An annual assessment plan is required to implement the ongoing evaluation of all intended sites.

### Pre-Assessment Steps

### (Complete these in order, refer to Appendix 1 for a pre-assessment checklist job aid)

- Assessors should be previously trained and competent in the use of the checklist and in conducting assessments.

- Prior to the visit, assessors should review the national testing guidelines for the pathogen testing programs included in the assessment and relevant (e.g., previous assessment reports) to determine whether previously identified gaps have been addressed.

- Assessors should work with MoH to identify and prioritize sites for evaluation in alignment with the MoH annual assessment plan.

- Agree on assessment dates with each site. Send a formal authorization to the facility stating the purpose of the assessment and the planned date and time. Because letters may be received late, follow up with a phone call a couple of days prior to the visit.

- Print at least one hard copy of the checklist for each testing site to be assessed. Where multiple testing sites exist within a single facility, each selected testing site must have its own completed checklist.

- At least one week prior to the assessment, send an electronic version of Part A (General Facility Information) of the checklist and request digital documents in advance.

- Arrange transportation and, where necessary, accommodation in advance of the assessment.

### Assessment

(Follow these practices during the site visit)

Introduction & Planning

- Introduce the assessment team and explain the purpose. Meet the site supervisor, testers, and health team. If a courtesy visit to the Regional or District Health Management Team is required, plan for it before the site assessment. Emphasize that evaluation is supportive and developmental, not punitive.

- Agree on the schedule. Confirm start and end times and build in reasonable breaks (tea and lunch) so the assessment proceeds smoothly and respects staff availability.

Review & Verification

- Review past gaps and improvements. Discuss any previously identified gaps and remedial actions the site has taken. Acknowledge completed improvements and encourage continued attention to any remaining recommendations. Maintain a constructive, supportive attitude.

- Verify facility/site details. Confirm the testing facility’s official name and physical address (as shown on official documents or facility entrance sign) and record dates in the country’s preferred format (see Part A).

Conducting the Assessment

- Use a calm and professional approach. Conduct the assessment respectfully. If a staff member does not understand a question, rephrase it clearly. Wherever possible, assessors should communicate in the site’s native language.

- Require evidence for verbal responses. Ask staff to back up verbal answers with documentation, current Standard Operating Procedures (SOPs), records, or examples of best practices.

- Confirm compliance by observation when appropriate. Direct observation by the assessor (for example, the presence and use of safety equipment, or watching the tester perform procedures) is considered best practice.

### Post-Assessment

(Complete these after the site visit)

- Debrief with the site team. Meet at the agreed time and location with testers, supervisors, and/or management to share main findings and recommendations. Be encouraging, highlight strengths and best practices, and frame recommendations as constructive next steps.

- Agree on the report timeline. Confirm the deadline for issuing the site assessment report and timeline required for remedial actions.

- Follow filing procedures. File the original site assessment report in accordance with MoH procedures.

- Deliver the report to the site. Ensure the site receives a copy of the report within the agreed timeframe.

- Offer follow-up support. Be available to site personnel who have questions or need clarification while implementing improvements.

- Invite feedback. Thank them for their time and ask any questions or suggestions to improve future assessments and the overall testing program.

### Tips on How to Conduct the Assessment

General Principles

- Use three sources of evidence. Evidence of compliance may be obtained through a combination of document review, observation, and staff interviews. For some requirements, direct observation alone may be sufficient (e.g., PPE use, workflow, hand hygiene). For others, documented evidence is essential (e.g., training records, QC logs, EQA reports).

- Apply professional judgment. Assessors should use professional judgment while ensuring that verbal responses are supported by observable or documented evidence wherever possible. Do not accept a "Yes" statement as valid unless confirmed by actual evidence (e.g., documentation or best practices).

- Ask open-ended questions. Use phrases like “show me how…” or “tell me about…” It is often unnecessary to ask all checklist questions verbatim. An experienced assessor can answer multiple checklist questions by asking open-ended questions of the testing staff.

Document Review

- Verify that written procedures exist and are current. Review all available records, logs, reports, standard operating procedures (SOPs) and/or job aids) to ensure they are available and current. Observe the written SOP/job aid, as testers might be able to explain the procedure, but it is not a written SOP/job aid.

- Check conformity with national guidance. Verify the correct algorithm and use of nationally approved tests (e.g., if a three-test algorithm is required for HIV diagnosis, confirm that the national algorithm is followed and only nationally approved tests are used).

- Follow up on prior findings. Confirm that the testing procedures are appropriate for the pathogen being tested and must follow the national guidelines. For example, if a three-test algorithm is required for HIV diagnosis, verify that the national algorithm is followed and that only nationally approved tests are used.

- Observe evidence that previous gaps were investigated, documented, and closed. If not closed, record corrective actions and add unresolved items to the report. If remediation is the responsibility of regional/national authorities, document and escalate to the designated point of contact.

Observation

- Observe the full testing sequence when feasible. Wherever possible, request permission to watch one complete test (pre-test, test, post-test) to confirm that practice matches written procedures.

Scoring Guidelines

- Yes (Y)= Full compliance; all elements are present and consistently implemented

- Partial (P)= Some aspects are in place but not completed to full conformance or not consistently implemented.

- No (N)= Lack of implementation in a particular area

- Non-applicable (N/A) = The question does not apply to the testing site or the testing program. This may be an option in a few cases. No score should be recorded. Rather, N/A should prompt a short explanation in the comments section.

Special Considerations

- Multiple testing sites. This checklist will be used to assess POC testing at a site(s) in a facility. Where a facility has numerous testing sites, it may not be feasible to assess each one during a single visit. It is therefore recommended to assess a subset of the testing sites (e.g., 3 or as directed by the per-national program) per visit, while ensuring that all rapid tests for all pathogens are included in the assessment.

- Logical scoring sequences. Take note in Section 4: each test might give different answers to the same question. For example, question 4.22” Is a standardized logbook/register available and correctly completed?”  might be Yes for HIV, No for Malaria, and Partial for Syphilis. Do not assume each test will score the same per question.

- Differentiate responsibility levels. It is important to differentiate the responsibilities of the testing sites per facility from the responsibilities of the regional or national level. Deficits in the latter are likely to result in a reduced overall score but highlight areas where the higher-level systems play a critical role in supporting and strengthening lower-level POC testing sites.

## Structure of the SPI-RDT Checklist

The SPI-RDT checklist contains four distinct parts:

- Part A: General Facility Information

- Part B: SPI-RDT Checklist:

- Organization and Management

- Physical Facility and Equipment

- Safety

- Testing

- Specimen Referral

- Part C: Scoring

- Part D: Site Assessment Report

### Part A.  Characteristics of the facility and testing site assessed

- Part A is a summary of general information on the testing site to be assessed. Limited facility-level information is collected only to provide context for the testing site and does not constitute a facility assessment.

- It is important to ensure that all areas are completed accurately.

- Note, there is no scoring for Part A.

Date of Assessment: Enter the date when the assessment is conducted using the European date format or the date as applicable in your country (e.g. dd/mm/yyyy). Note that the date format must be agreed upon during tool customization.

Time of Assessment: Enter the time when the assessment started using the format applicable in your country (e.g. 24hrs or 12hrs (AM/PM). Note that the time format must be agreed upon during tool customization.

Date of previous assessment: Enter the date of the previous assessment if available using the European date format or the date as applicable in your country (e.g. dd/mm/yyyy).

Testing Facility Name: Enter and accurately spell “Testing Facility Name” as indicated on official documents or facility entrance sign, e.g., Makonya Health Center.

Testing Site Name: Enter the name of the Testing Site to be assessed (e.g., TB clinic, OPD, ANC, Maternity, etc.).

Physical address or location of the testing facility: Provide the physical address/location of the facility.

Number of POC testing sites within the same facility (contextual information only): Indicate the total number of POC testing sites located within the same facility as this testing site, including those not assessed during this visit. This information is collected for contextual and planning purposes only and does not imply a facility-level assessment.

List of POC tests conducted at this testing site: Provide the list of all rapid diagnostic tests used.

Type of Facility: Check one which is most applicable

Level: Confirm with the supervisor about the level of the facility within the tiered health network and check the correct box. Note that as part of the customization, levels should be updated to reflect the local context.

Affiliation: Confirm with the supervisor about the affiliation of the facility and check the correct box.

Name and Title of staff: Provide the full name and job/position title of all staff performing POC testing.

Name of Interviewee(s), Title and Phone number: Provide the full name of job/position title and the phone number of the person being interviewed (or answering the questions) for all testing-related questions, ideally the actual tester instead of the supervisor.

| Date of Assessment: Testing Facility Name: | Time of Assessment: Testing Site Name: |  | Date of Previous Assessment: |
| --- | --- | --- | --- |
| Physical address or location of the testing facility: Number of POC testing sites within the same facility (contextual information only): Name of testing site(s) location(s) assessed: |  |  |  |
| List of Point of care (POC) tests conducted at this testing site: |  |  |  |
| Type of Facility (Check one which is most applicable) - Health Center ☐ Mobile Clinic ☐ Laboratory (Specify): - Specialized Clinic (Specify): ☐ Hospital Bedside ☐ Other (Specify): |  |  |  |
| Level (Check one) - Region, province or zone - District - Health center - Other (Specify): |  | Affiliation (Check one) - Government - Private - Faith-based Organization - Non-governmental organization - Other (Specify): |  |
| Name of staff |  | Title of staff |  |
| Name of interviewee: Title: Phone number: |  |  |  |

### Part B. SPI-RDT Checklist

Part B constitutes the main body of the SPI-RDT Checklist for conducting the inspection and scoring. The Checklist is laid out in five sections that align with the national standard requirements.

Instructions as per guidelines

- For each question, answer as follows:

- Tick Yes (Y), only if full compliance is demonstrated

- Tick Partial (P), if some compliance is demonstrated

- Tick No (N), if there is no evidence of compliance

- Tick Not Applicable (N/A), if the question is not applicable or irrelevant to the country context, a score is not required

- In the Comment box, describe the specific gap (Partial and No scores) or the reason the question does not apply (N/A)

- Sections 1 (Organization and Management), 2 (Physical Facility) and 3 (Safety) pertain to the overall facility incorporating the testing sites. Section 4 (Testing) assesses the testing site(s).

- ORGANIZATION AND MANAGEMENT

- This area assesses whether the testing site is operating in alignment with national guidelines

- It is important that the testing site follows written national guidelines, such as POC Testing Site Guidelines. Ideally, a copy should be available at the facility.

- Observe organogram/structure/ POC guidelines

- Determine if the testing sites are periodically supervised, and whether assessments are conducted

- Determine if gaps and recommendations are addressed

- Assesses processes e.g., working rosters, and necessary documents such as job descriptions are available

|  | 1.0 ORGANISATION AND MANAGEMENT |  |
| --- | --- | --- |
| 1.1 |  | Are lines of authority and roles and responsibilities clearly defined and an organogram available? Personnel should have a basic understanding of the organization they work for. A facility organogram should be available and displayed for all personnel to see. The organogram describes the lines of authority at the facility. Check evidence of organizational structure. As testers, do they know who to report to and who does their supervisor report to? Y = Evidence of the internal facility organogram, and personnel understand who they report to at the facility level P = Personnel have some knowledge about their organization and who they report to at the facility level, but there is no organogram, or there is an organogram, but personnel do not understand it. N = There is no evidence of an organogram, and personnel have no understanding of the facility's general structure, roles, and responsibilities, or where the facility fits into the general MoH structure. |
| 1.2 |  | Does each staff member involved in testing have a current job description? Each tester should have a written job description that is current. The job description should provide clear duties and tasks conducted by the tester. Ask the testers to describe their tasks and cross-check against the job description. The job description should be current and reflect the personnel duties. In cases where nurses are conducting rapid testing, this task should be included in their job description. Y = Job descriptions are available and current for all testers P = Job descriptions are available for some testers, or outdated or inaccurate N = No evidence of job descriptions |
| 1.3 |  | Where there is more than one tester, is there a duty roster? If there is more than one tester at the testing site being assessed, look for evidence of rotation among testers and review the weekly/monthly duty roster (e.g., 1 testing site + multiple testers). If there are several testers within the testing site, look for evidence of rotation to the various testing sites and review the weekly/monthly duty roster (e.g., multiple testing sites + multiple testers) Y = Evidence of a duty roster and rotation of staff, current and displayed P = Evidence of an old or outdated duty roster or no rotation of staff N = No evidence of a duty roster N/A = The site only has one tester |
| 1.4 |  | Are there adequate staff for the testing volume? Do testers complete the testing of all patients/specimens each day, or are some patients asked to return the following day? Y = All POC testing is completed the same day P = Some POC testing is completed same day, but other tests are not completed N = None of the POC tests are completed on the same day |
| 1.5 |  | Are minutes of routine (at least monthly) staff meetings available? Staff meetings are an important form of communication between staff. How often meetings are held depends on the site/facility, but generally, they should be held at least monthly. Observe evidence of meeting minutes and completed follow-up actions. The meeting could be between the tester(s) and supervisor(s), or the tester and the facility manager. This question aims to ensure regular communication. Y = Evidence of ongoing regular staff meeting minutes, including follow-up actions addressed for issues reported P = Evidence of occurrence of sporadic staff meetings, but there are follow-up actions on issues communicated by testers N = No evidence of any meeting minutes/communication |
| 1.6 |  | Are training records available on general basic quality activities for staff performing point of care testing? All personnel performing POC testing should have a basic understanding of a quality management system (QMS), including safety, record keeping, stock management, equipment (e.g., fridge, pipette), PT/EQA and QC, and testing procedures. Observe tester training records, logs, and/or certificates for QMS training. There should be a folder/log/register where all personnel records are stored. If quality management systems may be confusing, ask what aspects of quality the staff was trained in, and if there is a need to paraphrase the question Y = All personnel have been trained in basic QMS, and training records are available P = Some personnel have been trained or understand what quality means regarding POCT, and/or training records are not available for those trained N = Personnel have not been trained or there is no evidence of training records |
| 1.7 |  | Are copies of previous assessment reports available? POCT sites should be assessed on a regular basis in accordance with POCT guidelines. Repeated use of the same POC testing site checklist will assist in quantitative monitoring the improvements of facilities over time. A copy of the last assessment should be available as well as an accompanying site assessment report. Y = Recent site assessment report available P = Old/outdated report available N = No reports available N/A = New site with no previous assessment |
| 1.8 |  | Have gaps and recommendations from the last assessment been addressed? The site assessment report clearly states the gaps, recommended actions, who is responsible for addressing the gaps, and when and by whom they were addressed. Are actions being taken to address gaps? Are action items completed? Note, some of the actions might be beyond the scope of the site or facility. Some actions might be the responsibility of the national or regional level. Assessors should follow up with the national level regarding gaps observed across the facilities that are the responsibility of the national level. Y = Written evidence of action items addressed P = Action items are addressed but not documented consistently N = No evidence of actions implemented and no documentation N/A = New site with no previous assessment |

- PHYSICAL FACILITY AND EQUIPMENT

- This area assesses the physical condition of the testing site

- The physical structure should be in good condition with adequate testing space

- Adequate power and clean, running water should be consistently available

- Testing should be conducted in a private area, away from other patients, e.g., at a health center in a private   room or segregated area; at the bedside with a curtain surrounding the area of the bed, patient, and testing process

- Adequate lighting (natural or electric) is needed for sample collection, testing, and reading results.

- Most of this section is completed through direct observation by the assessor

|  | 2.0 PHYSICAL FACILITY AND EQUIPMENT |
| --- | --- |
| 2.1 | Is there a designated (private if needed) specimen/patient collection/reception area where appropriate? There should be a designated area to receive patients or specimens. This could be a waiting area outside the testing area for clients. Where specimens are received for testing or referral, is there a designated area for specimen reception? If the site tests clients only, then there needs to be a waiting area for them. Y = Designated area for patients, or specimen reception (if appropriate) P = Some testing sites have waiting areas for clients, but not all, or specimen reception is required, but is not in place (if appropriate) N = No designated area for patients or specimens |
| 2.2 | Is the testing area access-controlled and secure? Testing areas need to be private and security-controlled The testing area should be separated from patients waiting to be tested, at least by a screen, but preferably in separate rooms. The testing area/room should have an appropriate sign reading “restricted access.” If the area is a room, it should be locked when in use. If testing occurs at the bedside, a curtain should be used for privacy. Y = Evidence of signage, and the door is closed and locked when in use, or a curtain is present at bedside P = Testing room door has signage or is locked, but not both N = No signage or locks, or bedside curtain |
| 2.3 | Is there an area for staff to store food and personal items, separate from the testing area? Staff should have a separate area for eating, drinking, and storing personal belongings, outside or away from the testing area. Y = There is a dedicated space/room for eating, drinking and storing personal belongings P = There is a dedicated or separate space for eating, drinking and storing personal belongings in the room where testing is occurring N = There is no dedicated or separate space/room for eating, drinking and/or storing personal belongings |
| 2.4 | Is there a designated testing area for testing only, separate from the “clean” administrative area? The area/room should be partitioned or clearly separated for testing (contaminated) and record keeping (clean) for paperwork/records/reports. Y = Evidence of segregated contaminated and non-contaminated spaces. This could be as simple as having 2 desks or tables in one room. P = Segregated areas are available, but not used N = Testing, reporting, logbooks, everything happens in one place e.g., one desk/table for all work |
| 2.5 | Is the testing area clean and organized? The testing area should contain only testing-related items, and the area should be neat, organized and uncluttered. Even at the bedside, items should be available in a neat and organized manner. Y = Testing area is neat and consists of testing items only P = The area is neat, but includes testing and non-testing items N = The area is cluttered and untidy |
| 2.6 | Is sufficient lighting available in the designated testing area? The designated testing area should have adequate natural or electric lighting to support safe specimen collection and accurate reading of test results. Y = Adequate lighting (natural and/or electric) is consistently available in the designated testing area during testing hours, and results can be read accurately without relocating or pausing testing P = Lighting is adequate only when electric power is available and the site has a defined, acceptable contingency during outages (e.g., testing is paused until lighting is restored or testing moves to a pre-identified alternative area that meets testing area requirements) N = Lighting is inadequate for specimen collection and reading results, or outages lead to testing continuing in poorly lit/non-approved areas, or there is no acceptable contingency plan |
| 2.7 | Does the site have adequate and appropriate storage for test kits, consumables and ancillary equipment? There should be a designated area(s) for adequate and appropriate storage of test kits and consumables. Test kits should be stored following the manufacturer’s instructions, away from direct sunlight and in a secure area. Consider also checking the facility's storeroom or pharmacy if the testing site does not have a dedicated storage space or is supplied daily. Y = All test kits and consumables are stored adequately and appropriately according to the manufacturers’ instructions; protected from sunlight; secure; organized; First Expired First Out (FEFO)/expiry management evident; no signs of damage/exposure) P = Storage exists, but one or more requirements are not consistently met, or some items are stored appropriately while others are not. Specify gaps observed e.g., temperature monitoring absent, items exposed to sunlight/heat, mixed with chemicals/cleaning agents, poor security, crowding, humidity/water risk, expired kits present, FEFO not practiced N = Storage is not adequate to maintain kit integrity, e.g., no designated storage, or conditions are clearly non-compliant/unsafe. Unsecured, exposed to heat/sunlight/moisture/contaminants, no ability to maintain required conditions, frequent stock left in open areas, such that kit quality cannot be assured. |
| 2.8 | Is there a current inventory SOP/job aid available and a system in place for ordering and managing reagents and consumables? Each site needs a SOP/job aid on how to order kits and supplies from the facility’s storeroom or pharmacy, the district laboratory, or wherever supplies are obtained (storage locations may differ by country). It is also important to understand the facility's process for ordering test kits and supplies; this should include who orders, when they order, what the timeline is and how stocks are inventoried. Observe any written procedures regarding ordering. Observe evidence of a stock log/book/card being used as part of inventory management. Y = Inventory SOP/job aid available and current and stock log maintained P = Either SOP/job aid or stock log are available, but not consistently maintained N = No evidence of SOP/job aid or stock log |
| 2.9 | Are temperature logs available and updated daily for storage as per SOP/job aid? The temperatures of the fridge, storeroom or facility pharmacy should be monitored daily and records maintained. Observe if temperature logs have been completed and maintained daily. Y = Storeroom and/or fridge temperatures (if the site has a fridge for reagent storage) are monitored daily, and records are maintained as per SOP/job aid P = Room and/or fridge temperatures are monitored but inconsistently N = No evidence of temperature monitoring / no thermometers available, and no SOP/job aid |
| 2.10 | Is the appropriate equipment available, calibrated, including timers, thermometers and other small equipment where applicable, and are records available? All small equipment should be calibrated and serviced as per the quality manual or national guidelines. Thermometers are required to monitor the temperature of storage and testing rooms according to test kit temperature ranges. The calibration process should be supported by the national entity overseeing the facility where the POC testing is conducted or the structure/department in charge of servicing health equipment. Y = Calibration of small equipment is routinely performed and documented P = Some small equipment is calibrated documented or inconsistently documented N = No evidence of any calibration of small equipment |

3.0 SAFETY

- This area assesses safety compliance at the testing site

- This includes the availability of safety equipment and adherence to safety measures by personnel

- Observe evidence of safety equipment, e.g. gloves, coats/aprons, segregated waste, handwashing station, etc.

- Observe compliance with safety measures by personnel

- Confirm written and current SOPs/job aids on safety are available as appropriate

|  | 3.0 SAFETY |
| --- | --- |
| 3.1 | Are there current general SOPs/ job aids in place to implement safety practices (e.g., use of aprons/coats, gloves, hand washing, waste management, etc.)? Review safety SOPs for accuracy. Check for job aids displayed (e.g., handwashing, PPE, PEP, waste management, etc.). Y = Safety SOPs/job aids are available and current P = Safety SOPs/job aids are available, but outdated N = No evidence of safety practices, SOPs or job aids |
| 3.2 | Are current SOPs and/or job aids available to address accidental exposure to potential pathogens (e.g., spills, needle sticks, injuries, or cuts)? Observe evidence of SOP/job aid specifically on how to address spillage and needle sticks, including a job aid on how to make diluted bleach for decontamination of spillage. Job aids, such as Post-Exposure Prophylaxis (PEP) instructions, should be posted. A current SOP that addresses all responses to accidents should also be available. Y = SOPs/job aids are available for how to address spillages and PEP, and how to make diluted bleach for disinfecting work surfaces and spills P = Some SOPs/job aids are available, but not all, or SOP/job aids that address accidental exposure to potential pathogens are available but outdated N = No evidence of SOPs or job aids for spills or needlestick injuries |
| 3.3 | Is personal protective equipment (PPE) available to testers in the testing area? Verify that an adequate supply of appropriate PPE is present and accessible to testers (at a minimum, gloves and coat/apron) to support uninterrupted testing. Y = Appropriate PPE is available and readily accessible to testers at the testing area (e.g., adequate supply of gloves in relevant sizes and coats/aprons available) P = PPE is available but not consistently adequate or accessible (e.g., limited quantities, intermittent stock-outs, available only in the storeroom/not in the testing area, missing components such as gloves or coats/aprons, wrong sizes). Specify what is missing/limited. N = Appropriate PPE is not available or not accessible to testers (e.g., no gloves and/or no coat/apron available for use) |
| 3.4 | Do all testers properly and consistently use PPE throughout the testing process? Observe testers wearing the appropriate PPE designated by national guidelines (e.g., use of gloves, aprons, lab coats, etc., during testing). Used gloves are discarded after testing and placed in the contaminated waste bin. Y = All testers wear appropriate PPE (e.g., gloves and aprons) for testing and gloves are appropriately discarded after testing P = PPE (gloves, aprons) are not consistently worn by all testers N = Testers do not wear PPE during testing or gloves are not discarded when testing clients has ended, or PPE is worn in the “clean” non-contaminated working area |
| 3.5 | Is clean water, soap (or hand sanitizer) and paper/hand towel available for hand washing, and being used by all testers per SOP/job aid? Observe evidence of a hand-washing area with water, soap, and paper/hand towel. In some instances, running water might not be available. In such cases, testers can use a bucket of water and jug to wash their hands. Ask the tester to wash his/her hands and assess adherence to the SOP/Job aid. Y = Water and washing supplies are available and consistently used by all testers to wash their hands in accordance with SOP/Job aid P = Not all handwashing supplies are available, or not all testers follow the SOP/job aid for handwashing N = No handwashing supplies available |
| 3.6 | Is the disinfectant solution available properly labeled with content, date of preparation, and date of expiration? Verify that the site has appropriate disinfectants. Check if the disinfectant container is labeled with the contents, date of preparation and expiration. Y = Bleach/disinfectant is freshly made and labeled correctly P = Bleach/disinfectant is available, but not labeled or labeled but expired N = Bleach/disinfectant container is not labeled or not available |
| 3.7 | Are sharps, infectious, and non-infectious waste separated in accordance with current waste management SOPs? Observe segregated waste (infectious, non-infectious, and sharps containers). The waste containers are either color-coded or labeled. Y = Waste is properly segregated and stored in accordance with the available SOP/job aid P = The waste is inconsistently or inappropriately segregated (e.g. sharp container filled beyond the limit, a sharp container properly used but contaminated waste mixed with non-contaminated waste) N = No evidence of waste segregation |
| 3.8 | Are infectious waste and sharps containers transported and disposed of safely? Ask the tester to describe the process for removing and disposing of the infectious waste and the sharps containers. Observe whether the waste is securely stored temporarily before pick-up and/or incinerated. Ask to see where the incinerator area is located. Y = Evidence of proper handling of infectious wastes and sharp containers (e.g., temporary storage, safe transport and disposal) P = Infectious waste and sharp containers are not stored appropriately and/or disposed of per the SOP/job aid N = No evidence of proper temporary storage or safe removal/disposal of infectious waste and sharp containers |
| 3.9 | Is vaccination against Hepatitis B infection available to testers, and is there documented evidence of vaccination or refusal? Personnel who work with whole blood or serum should be vaccinated in accordance with national health policies and guidelines if available. Y = Personnel are vaccinated, and records are available P = Some, but not all, personnel who handle specimens are vaccinated and records available, or personnel are vaccinated but no records are available N = No evidence of vaccination or vaccination refusal N/A = If (no national requirements for Hep B vaccination, this is question is not applicable) |

- TESTING

- In this section, each diagnostic test or testing algorithm (e.g. HIV) per pathogen is assessed.

- This area covers all aspects of testing, and all three phases (pre-test, test and post-test) are adequately implemented in accordance with SOPs/job aids.

- Basic Quality Management Systems (QMS) are important at testing sites and throughout the facility. Quality measures and appropriate implementation of QMS are essential to ensure quality test results.

- Testers should be trained in basic QMS training that will crosscut all tests.

- Directly observe to assess pre-test, test and post-test processes and adherence to SOPs. Be specific about which processes are deficient.

- Observe all QC and EQA/PT records. The testing sites should conduct EQA/PT/retesting/blinded testing according to national guidelines for each test.

- SOPs/job aids should be available at the testing sites within a facility.

- Check that all SOPs/job aids are current and available for all tests and processes conducted.

- Observe the tester performing testing wherever possible or request a mock demonstration of the testing.

- Check that training and competency records are available and up to date.

- Examine all records and logs for complete and accurate information or test results.

- Ensure all other related documents are available and current.

| 4.0 TESTING |  |  |
| --- | --- | --- |
| Name of pathogen tested e.g. HIV |  |
| Manufacturer name of test. Name all tests where there are testing algorithms. e.g. HIV First Complete for each test or algorithm – see example in Appendix 2 |  |
| 4.1 | Are testers initially trained in-service, deemed competent, and are records available? Testers are required to receive in-service training and be deemed competent. Certificates of training and competence should be available. Ask for personnel files or personnel training records/logs to confirm training was completed. Y = Evidence that all testers at the site have a record of training in testing procedures (e.g., training certificate, training log attendance sheet, etc.) P = Evidence that some testers at the site have a record of training in testing procedures (e.g., training certificate, training log attendance sheet, etc.) N = No evidence the testers at site have a record of training in testing procedures (e.g., training certificate, attendance sheet, etc.) |
| 4.2 | Are records of periodic competency assessments conducted for all testers? Observe evidence of periodic competency assessments and records for all testers. Ask for the personnel file, personnel training records, or log to confirm that periodic competency assessments occur. Y = All testers have records of recent competency assessments P = Some testers have records of recent competency assessments, or the assessments are not current N = None of the testers have competency assessments or records |
| 4.3 | Is the test conducted nationally approved as per SOP/job aid? Each test conducted should be an approved test and follow the per SOP/job aid. Observe that the tests being conducted are the same tests as per SOP/job aid. Y = Evidence that all tests used are nationally approved P = Evidence that some tests used are nationally approved (e.g., Two of three HIV tests are approved, but the third is not included in the national testing algorithm). For single test algorithms, evidence of inconsistent use of nationally approved test N = No evidence of use of nationally approved tests |
| 4.4 | Is the testing SOP/job aid available and current? The POCT sites are not required to write their own SOPs and/or job aids. These should be readily available from the supporting level above or the regional/national levels. If SOPs and/or job aids are outdated according to the quality manual or national guidelines, the facility manager should request updated SOPs and/or job aids. The testers must read the SOPs/job aids, and records of reading and comprehension must be available. Y = The SOPs/job aid is current and available, and there is evidence that testers have read and understand the SOP. P = The SOP or job aid is outdated and/or there is no evidence that testers have read the SOP/job aid. N = No SOP nor job aid available |
| 4.5 | Is the job aid displayed or accessible? Job aids should be available/provided at the national/regional laboratory level, posted where testers perform testing, and made available for all tests. Y = The job is available and displayed at the testing site P = The job aid is available, but is torn or difficult to read, or not posted N = No job aid is available |
| 4.6 | Is Quality Control (QC) testing performed routinely as per SOP? All tests must use quality control materials in accordance with the national guidelines. The type of QC and how often they are run will depend on the type of test and the availability of QC. This question depends on the test and the SOP/guidelines. For some tests, QC might be run daily, while others might be run weekly or monthly. Y = There is evidence of QC being run and recorded in accordance with SOP/guidelines/quality manual P = QC is run irregularly and/or not recorded consistently N = No evidence of QC being run |
| 4.7 | Are QC results documented and reviewed by the tester prior to reporting patient results? Check that QC results are properly documented in the testing register or QC log as per SOP and/or job aid. Is there evidence that the QC results are reviewed by the supervisor or the lab technician/technologist? Check that failed QC is also recorded, and corrective actions are documented. Y = QC results are run and recorded consistently and correctly. Evidence that QC results are reviewed, failed QC runs are investigated, and corrective actions documented P = Inconsistent documentation of QC runs or review of QC results, or lack of follow-up actions for failed QC runs N = No documentation of evidence of QC runs nor review of results |
| 4.8 | Is the site enrolled in an External Quality Assessment (EQA) program (e.g., Proficiency Testing (PT), re-testing or inter-facility comparison) for the test performed? The national or regional/district provides PT panels to laboratories and testing sites per the national guidelines. Sites not enrolled in EQA are scored zero (N), and this might highlight gaps at higher levels within the tiered laboratory network. If available, PT panels should be provided to all sites and laboratories for the diagnostic tests performed. However, when no PT panel is available, verify that the laboratory or testing site is implementing another method of EQA (re-testing/rechecking or inter-site comparison). Y = Evidence that the site consistently participates in any EQA program. Check the most recent EQA cycle. PT panels are provided for the POC test performed, or any other EQA method is conducted (state which type of EQA is used in the Comments section) in accordance with national guidelines/policy P = Evidence of inconsistent participation in an EQA program: PT panels distributed to the POC site is irregular or there is a lack of alternative EQA methods N = No evidence of participation in any EQA program |
| 4.9 | Are EQA reports (e.g. PT, re-testing, inter-facility comparison) reviewed and is there recorded evidence of corrective actions for failed EQA results? Confirm that, regardless of the EQA method used, the reports are available, reviewed, and appropriate corrective actions are implemented and recorded for failed EQA results. Y = The site participates in an EQA program for the test performed, the EQA reports are available and reviewed, and there is evidence of corrective actions taken for failed EQA results P = The site participates in an EQA program for the test performed and the EQA reports are available, but there is no evidence of corrective actions taken for failed EQA results N = No evidence of EQA reports reviewed, nor corrective action taken for failed EQA results |
| 4.10 | Does the site have appropriate equipment for the test performed? In some settings, the site may use some equipment (e.g., timers, pipettes, thermometers, fridges, centrifuges). Y = Necessary equipment is available for the test. Thermometers should be available to monitor the temperature of test kits. P = Some, but not all, necessary equipment is available for the test N = The necessary equipment is not available for the test N/A = No equipment required |
| 4.11 | Are test kit inserts available for all tests at the testing site? Kit inserts should be available for testers to refer to as necessary, e.g., for lot numbers or additional information not provided in the job aid. Y = Kit inserts are available at the testing site P = Kit inserts are available for some but not all tests N = No evidence of any kit inserts |
| 4.12 | Are the test kits stored according to the manufacturer’s instructions? Observe how and where kits are stored. Kits should be well organized and stored, observing First Expired First Out (FEFO). Check that temperatures are suitable for the kit storage. For example, if kits should be stored between 20 – 30 degrees Celsius, but the room temperature is 35 degrees, this is not appropriate for storage. Y = Test kits are stored appropriately at room temperature or refrigerated at appropriate temperatures in an organized and logical FEFO. P = Test kits are stored under appropriate temperatures, although temperatures are not routinely monitored, or the kit storage is unorganized according to FEFO N = Tests are stored outside of the accepted range of temperatures |
| 4.13 | Are new batches/lots/shipments of kits verified prior to use and records available? It is important that new batches or shipments of kits be verified before testing client specimens. QCs should be included in the verification of the new batch/lot/shipment. Records of new batch results should be maintained and made available, e.g., by checking whether the current in-use batch has been verified and whether records are available. The instructions should be in an available SOP and/or job aid. Y = Evidence of QCs being used on new test batches and records available P = According to the tester, QCs are run on new batches/lots to verify the lot, but no records available N = No evidence of verification of new lots/batches using QC materials |
| 4.14 | Are adequate test kits available? Ask whether the site experienced any test kit stockouts and when the last one occurred. Review the testing register to confirm that no testing was performed during that period; review the monthly/quarterly statistics to calculate the number of tests conducted. Then check the stock cards or storage area to confirm that the site has an adequate stock level of kits for the month or quarter. Y = Adequate test kits are available P = Inadequate stock level of test kits N = There is a stockout of tests |
| 4.15 | Are adequate consumables available? Ask whether the site experienced any stockout of consumables needed to perform the test (e.g., lancets, needles, swabs, stationery, any non-test kit items) and when the last occurrence was. Review the monthly or quarterly statistics. Check the stock cards or storage area to confirm that the site has an adequate stock level of consumables for the month or quarter. Y = Adequate stock level of consumables P = Some consumables, but not all available (list in the comment section the consumables that are not available and since when) N = No consumables available to perform the test |
| 4.16 | Are expired kits safely disposed of or removed from the testing site as per SOP/job aid? No expired kits should be stored on-site. They should be disposed of/removed from the testing site in accordance with the national SOP/guidelines. Y = No expired kits are stored at the testing site nor used for patient testing as per SOP/guidelines, and the tester understands the facility-level procedure for handling expired test kits. P = Expired kits are stored at the site OR used for patient testing N = Expired kits are stored at the site AND used for patient testing |
| 4.17 | Is there a specific workflow for pre-testing, testing and post-testing? Observe the workflow of all three phases. Is there a logical and standardized workflow for pre-testing, testing and post-testing processes? Could the workflow be improved? Observe the process, or the tester can describe the process if observation isn’t possible. Y = Evidence of an optimal workflow for the test or the testing algorithm from sample collection to reporting P = Some areas of workflow are sub-optimal (results aren’t recorded immediately, e.g.) N = No evidence of optimal workflow. The process is not streamlined. Examples: - Testers testing more than one client at a time. This could lead to mismatched results or a breach of patient confidentiality. - Tester writes client details in the logbook after testing, instead of before testing - No waiting place for clients, or no logical queue of clients to be tested - Tester does not have all the reagents/supplies in hand before testing - Test results are written in the logbook after the client has left - Tester does not wash his/her hands after testing |
| 4.18 | Is testing conducted in accordance with the SOP/job aid? Observe, if possible, whether the tester is performing the test per the SOP and/or job aid. If no testing can be performed, ask the tester to walk the assessor through the procedure. Y = All testers conduct testing in accordance with a SOP/job aid P = Some testers conduct testing in accordance with SOPs/job aids, or evidence that the tester deviates from the procedure sometimes N = Evidence of deviation from the testing procedure as outlined in the SOP/job aid/kit insert |
| 4.19 | Are nationally approved algorithms used where applicable (e.g. HIV testing algorithm)? Observe, if possible, the tester is performing the test per the SOP and/or job aid. If no testing can be performed, ask the tester to walk the assessor through the procedure. Observe all testers for each specific test. Y = All testers conduct testing in accordance with a SOP/job aid P = Some testers conduct testing in accordance with SOPs/job aids, or evidence that the tester deviates from the procedure sometimes N = Evidence of deviation from the testing procedure as outlined in the SOP/job aid/kit insert |
| 4.20 | Are corrective measures recorded and documented, as per SOP/job aid, when a test result is invalid or incorrect? Occasionally, a test result might be invalid or incorrect. Invalid results may be caused by a defective test device (e.g., strip, cassette), tester error, expired kits, or sub-optimal storage conditions. On the other hand, incorrect test results may be due to specimen mix-up, a deviation from the testing procedure, or a clerical error. Therefore, it is important that testers identify the root cause, take and document the corrective action as per the SOP/job aid. Ask the tester if they have had invalid or incorrect results; if so, have them describe the process. Review the testing register for the past 3-6 months, depending on the volume tested, to identify any invalid or incorrect test results and the related corrective actions taken. Y = For invalid and incorrect results, the site documented the corrective action implemented in accordance with the SOP P = Inconsistent documentation of invalid and incorrect test results and/or corrective actions records N = The site does not document invalid results or corrective actions |
| 4.21 | When errors/incidents or accidents occur (e.g. per site assessment report), are corrective measures recorded and documented as per SOP/job aid? Apart from invalid results, various other mistakes/errors can occur (e.g., incorrect result reported, needle stick, failed PT etc.) and these errors need to be resolved and records maintained. Y = Corrective action records are available and carried out in accordance with the SOP/job aid P = Corrective action records are available, but no SOP/job aid available or partial/incomplete documentation, or SOP/job aid is available, but corrective actions are not taken N = No evidence of corrective actions recorded and no SOP available |
| 4.22 | Is a standardized logbook/register available and correctly completed? All test results should be properly documented, and each testing site should use standardized logbooks/registers to record the testing results, which should be available for all testing. Review that the logbook or testing register for accuracy and completeness is correctly completed. If an electronic system is used, check the system's records for completeness and accuracy. Y = Standardized logbooks/testing registers are available and have been correctly completed for all fields P = Standardized logbooks/testing registers are available, but are incomplete N = No standardized logbook/testing register used by the testing site |
| 4.23 | Are completed logbooks/registers properly archived? Completed logbooks/registers need to be stored neatly, labeled, and safely. The site should have a system for archiving all completed logbooks. If an electronic system is used, check the procedure for correct and consistent backup. Y = Completed logbooks/registers are labeled, stored properly, and tidy in a secure room P = Completed logbooks/registers are stored, but not properly archived or not in a clean, secure place N = Logbooks/registers are not archived |

Note: Ensure all questions are asked per test as appropriate. Do not assume that because something is present or absent in one test, it will be the same for all tests. This means the possible scores for each test might differ, as might the actual scores. Be mindful when scoring, especially in Section 4, but also for all sections.

4: Gaps or observations from Section 4 above

Since there is no space above, all questions that scored a P or N need a comment in the table below explaining the score obtained. Describe the gap identified for improvement. This helps provide specific gaps for each test question.

| Test Program Name | Question Number | Score (P or N) | Gaps/ Observations for improvement | Comments |
| --- | --- | --- | --- | --- |

5.0 SPECIMEN REFERRAL (to a higher level for specialized lab testing)

- Not all testing sites refer to specimens for further testing. If not, cross through the section and put non-applicable (N/A).

- Check for evidence of referral SOP, logbooks for tracking, etc.

- Observe adequate availability and use of triple packaging

- Check if drivers have been trained in safety and if training logs are available

| Note: If the testing site does not refer specimens, mark N/A for Section 5.0 and note “No specimen referral conducted at this site |  |
| --- | --- |
| 5.0 SPECIMEN REFERRAL |  |
| 5.1 | Is there a current SOP/job aid for referring specimens which includes collection, storage, transportation, tracking and reporting? If the site refers specimens to a referral site (e.g., dried blood spots for HIV Viral Load), a general SOP/job aid should be available that describes the step-by-step process of referring specimens (including specimen collection, storage, packaging and transportation). Y = SOP/job aid is available and current and describes the general referral procedure P = The SOP/job aid is available but outdated N = No SOP/job aid available for specimen referral |
| 5.2 | Are records and results of referred specimens available based on the current SOP/job aid? Records/logs and the monitoring of report receipt from the referred laboratory need to be updated and current. Expected turnaround times for each referred test should be available. The testers should have a process in place, as per SOP/job aid, to monitor the return of reports. If reports are delayed, they should be followed up by the tester, and all referred specimens and reports received should be maintained in a Referral Logbook. Y = Consistent and complete record keeping for all specimens sent for additional testing in a referral log per the SOP/job aid P = The site has records as per above, but they are not consistently maintained N = No records exist for referred specimens for testing |
| 5.3 | Are adequate packaging materials available and correctly used per SOP/job aid? Observe if the facility has the appropriate materials to package specimens safely, and if feasible, observe how they package specimens for shipping. Ideally, some form of triple packaging should be used. Do personnel understand triple packaging and is triple packaging used? At lower levels, expensive triple packaging might not be available, but there should be a procedure for safely packing specimens per SOP and/or job aid. Y = Adequate triple packaging available and correctly used P = Some triple packaging available or materials available, but incorrectly used N = Stock out of triple packaging materials |
| 5.4 | Are referral specimen transportation logs included with specimens for tracking? When specimens are transferred, a completed log describing the number of specimens referred and a list of identifying numbers should accompany them. This log is transported with the specimens and used at the referral laboratories to verify that the specimens referred are the same as those received. This information is also transferred into the on-site referral log. Y = A referral log form is completed and transported with the referred specimens P = A referral log form is incomplete, or not transported with the referred specimens N = No referral transport logs are used |
| 5.5 | Is a referral record/logbook available and maintained? This logbook should be available and maintained. The list of specimens being referred should be complete and dated upon receipt of reports. If reports are not received, the tester should follow up with the referral laboratory and the log updated until the reports are received. In some cases, referred specimens are logged into the routine client register/logbook. In such cases, check the register for complete information. Y = A referral logbook is available, complete and updated P = A referral logbook is available, but incomplete, or delayed reports lack follow-up N = No logbook is available. Referred specimens are not recorded. |
| 5.6 | Does the facility receive copies of referred test reports from the reference lab and are these distributed to the test requester? When reports are returned to the referral site, there should be logged evidence of when they were received and when they were sent out to the requester or as per Referral SOP and/or job aid. Y = POC site receives the result reports from the reference lab, updates the logbook, and sends results to the requester. P = POC site receives some results but does not update the logbook or does not report to requesters N = No system is in place for sending referred results to requesters |
| 5.7 | Does the testing site monitor the turnaround time (TAT) of referred specimens from the referring facility to receipt of reports from the reference laboratory? Turnaround times (set by the referral laboratory) should be available and posted at the testing site. Y = Turnaround times are monitored for all referral tests P = Turnaround times are monitored for some referral tests N = Turnaround times are not monitored for referral testing |
| 5.8 | Where the result reports are delayed, does the site follow up with the referral laboratory and document the action? If result reports are delayed, the tester should follow up with the referral laboratory, and records of the follow-up (e.g., phone call, email, SMS, etc.) should be maintained. If turnaround times become chronically delayed, this is an ongoing problem that needs to be resolved between the two labs and corrective actions recorded. Y = Where TAT is delayed, the facility has followed up and recorded communication P = Facility follows up on the delayed TAT of results, but there is no record N = Delayed TAT results are not followed up |
| 5.9 | Is there appropriate temporary storage space for referral specimens prior to transportation? Observe where the referral specimens are stored (e.g. refrigerator, ambient temperature, etc.) before being transported to the referral site. Are the space and temperature appropriate according to established SOPs/job aids, separated from reagents and testing consumables, and safe and secure until shipment? Y = All referral specimens are appropriately stored prior to transportation P = Not all specimens are stored correctly prior to transportation. (e.g., either the ambient temperature is too high, or there is no fridge/cooler for specimens requiring refrigeration N = Referred specimens are improperly stored |

### Part C. Scoring

This section outlines the scoring criteria and explains how to calculate actual scores and final percentages based on the total possible score. This section is completed post-assessment visit.

- YES indicates ALL elements are present. Evidence of compliance must be available through observation or the review of documents (e.g., SOPs, job aids, records, logs, etc.).

- PARTIAL indicates some implementation of compliance but not full or inconsistent compliance, e.g. Job descriptions are available for some but not all testers.

- NO indicates no evidence of compliance, e.g., no training records of staff performing testing, staff do not follow the SOP

- Non-applicable (N/A) means a question is not relevant to the POC testing site.  Be aware of such answers as they need to be eliminated from the total possible score.

- YES is awarded 2 points, PARTIAL is awarded 1 point and NO is awarded zero points

- Note: each section’s total possible score needs to be calculated by the assessor:

  - If an answer is N/A, that question must be excluded from the total possible score, e.g. if a section is worth 20 points (10 questions) and 9 questions score YES (2 points each), but one question is N/A, then the total for that section is 18 out of 18

  - In Section B4 Testing, the total possible score will depend on the number of tests conducted

  - Within Section B4 Testing, the total possible scores might vary significantly across the tests according to the quality measures implemented per test.

- After completing section B, calculate all possible scores and all actual scores obtained

- Convert the overall score to a percentage and rate the site according to the percentage ranges and levels below

- Where there are multiple testing sites per facility, it might be feasible to only assess a selection of testing sites. In such cases, ensure the testing sites assessed include all pathogen rapid tests conducted at the facility.

| Assessment Score Sheet Summary |  |  |  |
| --- | --- | --- | --- |
| Section B | Section Name | Total Possible Score | Total Score Obtained |
| Section B1 | Organization and Management |  |  |
| Section B2 | Physical Facility and Equipment |  |  |
| Section B3 | Safety |  |  |
| Section B4 | Testing |  |  |
| Section B5 | Specimen Referral |  |  |
| TOTAL SCORE |  |  |  |

This checklist consists of five different levels to indicate status toward national certification.

| Levels | % Score | Description of results |
| --- | --- | --- |
| Level 0 | Less than 40% | Needs improvement in all areas and immediate remediation |
| Level 1 | 40% - 59% | Needs improvement in specific areas |
| Level 2 | 60%-79% | Partially eligible for site certification |
| Level 3 | 80%-89% | Close to national site certification |
| Level 4 | 90% or higher | Eligible for national site certification |

Assessor 1 Name:							Assessor 2 Name:

Assessor 1 Signature: 					             Assessor 2 Signature:

Time at end of audit (HH:MM):

### Part D. Site Assessment Report and Recommendations

- The site assessment report is written and completed post-assessment visit

- This is based on the scores and comments included in the assessment section

- It is important that comments in the assessment are specific to avoid ambiguity

- For every PARTIAL or NO, a gap or observation identified for improvement will be written

- The gap or observation should be specific, easy to read and understand, so that site personnel are able to take appropriate remedial action

- Recommendations should be specific to resolve a particular gap and should be clearly written. This is also an opportunity to establish and highlight if the recommendation is within the control/scope of the testing site or needs to be addressed at a higher level within the MoH.

- A copy of the evaluation, scores and site assessment report should be submitted to the site within a 2-week period and the original archived as appropriate

- It is possible that the evaluation and site assessment report can be completed electronically

| SPI-RDT SITE ASSESSMENT REPORT |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Name of Site/Facility: |  |  |  |  | Date: |  |  |
| Question Number | Test Name | Score P/ N | Gap/observation for improvement | Recommendation | Site/Facility/District/ Regional or National Responsibility | Date Completed | Responsible Person |

Assessor(s) Name:	                                               Signature:	                                                               Date:

Appendix 1

PRE-ASSESSMENT PREPARATION CHECKLIST

| Item | Action to be taken | Date for completion (dd/mm/yyyy) | Responsible person | Completed (Yes or No) | Comments |
| --- | --- | --- | --- | --- | --- |
| - Identification of sites | XXX sites approved by MOH |  |  |  |  |
| - Notification of sites | Letters sent to XXX number of sites |  |  |  |  |
| - Send the electronic version (Part A general facility information and optional testing statistical summary) of checklist to the sites | XXX checklists sent to XXX number of sites |  |  |  |  |
| - Arrange logistics (transport and accommodation) | Transport and accommodation arranged |  |  |  |  |
| - Print one copy of checklist for XXX sites | Copies printed for XXX sites |  |  |  |  |
| - Review of previous site assessment report | Previous reports reviewed for XXX sites |  |  |  |  |
| 7. Prepare list of standard tools (e.g. register, logbook, manual, training records, etc.) and documents (e.g. SOP/job aids, sample transport form, etc.) to be reviewed during the site assessment |  |  |  |  |  |

Appendix 2: Completed checklist example

Appendix 3: Example of a completed site assessment report and recommendations